{"text_format":"3","workbook":"xls","n_sheets":1,"sheets":["输尿管"],"not_rendered":{"external_links":1},"read":"whole","date_system":"1900","formulas":"cached_only"}

{"sheet":"输尿管","index":0,"n_rows":43,"n_cols":4,"rows_kept":43,"merged":[[0,0,0,3],[1,1,1,3],[2,1,2,3],[3,0,3,3],[38,0,38,2],[39,0,39,2]],"flags":["formulas_cached_only"]}
["技术参数"]
["项目\n名称","电子输尿管肾镜系统"]
["预算单价（万元）",100]
["技术参数表"]
["序号","技术和性能参数名称","性能要求","备注"]
["★1","基本要求","具备NMPA认证，同时具备CE或FDA认证，用于泌尿系统的检查与治疗"]
[2,"摄像系统主机","高清医用内窥镜数字摄像系统"]
[2.1,"采集分辨率","≥1920*1080"]
[2.1,"主机扫描方式","逐行扫描技术"]
[2.2,"软件功能","具备静态、动态图像存储、输出、画中画显示等"]
[2.3,"接口","具备高清输出信号接口、DVI输出接口等"]
["▲2.4","色调调节","调节范围≥16档"]
["▲2.5","特殊光功能","具备特殊光电子染色技术"]
[3,"冷光源","具备"]
[4,"电子输尿管肾盂镜"]
[4.1,"观察范围","视野角≥90°"]
[4.2,"镜子直径及器械通道","先端部直径≤8.5Fr，插入部直径≤10.1Fr，器械通道≥3.6Fr"]
["▲4.3","弯曲角度范围上","≥180°"]
["▲4.4","弯曲角度范围下","≥275°"]
[4.5,"操作遥控按钮","≥2个"]
[4.6,"有效景深","景深2mm－50mm"]
[4.7,"长度","工作长度≥660mm，总长度≥980 mm"]
[5,"输尿管硬镜"]
["▲5.1","小儿硬性输尿管镜","先端部直径≤6.8Fr，插入部直径≤8.5Fr，器械通道≥4.1Fr"]
["▲5.2","成人硬性输尿管镜","先端部直径≤7.8Fr，插入部直径≤10.5Fr，器械通道≥6Fr"]
[5.3,"有效长度","≥410mm"]
[5.4,"灭菌方式","可高温高压、低温等离子、环氧乙烷气体灭菌"]
[6,"图文工作站"]
[6.1,"基本功能","图像采集、报告打印、视频录像等"]
[6.2,"视频剪辑功能","具备"]
[6.3,"影像采集","可同时采集动态影像及静态影像"]
[7,"高清液晶监示器"]
[7.1,"尺寸","≥27寸"]
[7.2,"分辨率","≥1920*1080"]
[8,"腔镜专用台车","具有阻尼式自由型调节支臂、具有棚板调节位，自由调节高度、监视器平台可上下调节"]
[9,"配置需求","摄像系统主机1套（含冷光源等附件）、电子输尿管肾盂镜1条、小儿、成人硬性输尿管镜各1条、监视器1台、台车1个、图文工作站1套、侧漏器1个（标准配置内的其它附件须单列注明）"]
[10,"质保","免费保修期(全保，包含人工费、零配件更换等各项费用)≥1年，供应商承诺在接到报修电话后，30分钟内电话响应，24小时内现场响应；"]
[11,"安装培训要求","现场免费安装培训"]
["注：1.在序号处可加“★”号，表明该条款均为实质性响应指标要求，必须全部响应。若有一项带“★”号条款未响应或不满足，均视为非实质性响应招标文件，按无效投标处理。\n    2. 在序号处可加“▲”号，表明该条款为重要指标，若满足，在评标中加分多于普通项。"]
[]
[" "]
[" "]
[" "]